Geochronological methods

Samples CD05 and CD06 were dated at the GeOHeLiS platform, University of Rennes (France) by laser ablation inductively coupled plasma mass spectrometry (LA-ICP-MS) using an Agilent 7850 Q-ICP-MS connected to an ESI Excimer Laser (NWR 193UC). The analysis consisted of 30 µm spot diameters produced with a repetition rate of 4 Hz and a fluence of 6.2 J/cm2 (Table 1). Ablated material was carried to the mass spectrometer in He (~0.77 l/min) and then mixed with N (0.05 l/min) and Ar (~0.70 l/min), before being introduced to the ICP source. Tuning of the instrument and mass calibration were performed before the analytical session using the NIST SRM 612 reference glass, by monitoring the 238U signal and minimizing the ThO+/Th+ ratio (<0.5 %). The analyses consisted in the acquisition of the 204(Pb + Hg), 206Pb, 207Pb, 208Pb, 232Th and 238U signals. The 235U abundance was calculated from the measured 238U on the basis of a 238U/235U ratio of 137.818 (Hiess et al., 2012). Single analyses consisted of ~20s of background integration with the laser off, followed by ~40s integration with the laser firing and then a ~10s delay for wash out. Raw data were corrected for Pb/U and Pb/Th laser-induced elemental fractionation and for instrumental mass discrimination by standard bracketing with repeated measurements of the zircon reference material GJ-1 (Jackson et al., 2004). Along with the unknowns, the 91500-zircon reference material (Wiedenbeck et al., 1995) was measured to monitor precision and accuracy of the analyses and produced a concordia age of 1065.8±3.4 Ma (MSWD=0.94, n=28) during the course of the analyses. Data reduction was carried out with the Iolite software package (Paton et al., 2010). The analyses were performed in time resolved mode. All the reference materials and data can be found in Supplementary File 4. More information about the analytical procedure can be found in the following Table and in Nosenzo et al. (2022).

Operating conditions for the LA-ICP-MS equipment

| Laboratory and Sample Preparation |  |
| --- | --- |
| Laboratory name | GeOHeLiS Analytical Platform, OSUR, Univ Rennes 1, France |
| Sample type/mineral | Zircon |
| Sample preparation | Epoxy mount |
| Imaging |  |
| Laser ablation system |  |
| Make, model and type | ESI NWR193UC, Excimer |
| Ablation cell | ESI NWR TwoVol2 |
| Laser wavelength | 193 nm |
| Pulse width | <5 ns |
| Fluence | 6.2 J/cm2 |
| Repetition rate | 4 Hz |
| Spot size | 30 μm |
| Sampling mode / pattern | Single spot |
| Carrier gas | 100% He, Ar make-up gas and N2 (3 ml/mn) combined using in-house smoothing device |
| Background collection | 20 seconds |
| Ablation duration | 40 seconds |
| Wash-out delay | 15 seconds |
| Cell carrier gas flow (He) | 0.77 l/min |
| ICP-MS Instrument |  |
| Make, model and type | Agilent 7850, Q-ICP-MS |
| Sample introduction | Via conventional tubing |
| RF power | 1150 W |
| Sampler, skimmer cones | Ni |
| Extraction lenses | X type |
| Make-up gas flow (Ar) | 0.70 l/min |
| Detection system | Single collector secondary electron multiplier |
| Data acquisition protocol | Time-resolved analysis |
| Scanning mode | Peak hopping, one point per peak |
| Detector mode | Pulse counting, dead time correction applied, and analog mode when signal intensity > ~ 106 cps |
| Masses measured | 204(Hg + Pb), 206Pb, 207Pb, 208Pb, 232Th, 238U |
| Integration time per peak | 10-30 ms (207Pb) |
| Sensitivity / Efficiency | 17000 cps/ppm Pb (50 µm, 10 Hz) |
| Data Processing |  |
| Gas blank | 20 seconds on-peak |
| Calibration strategy | GJ1 zircon standard used as primary reference material, 91500 used as secondary reference material (quality control) |
| Common-Pb correction, composition and uncertainty | No common-Pb correction. |
| Reference Material info | GJ1 (Jackson et al., 2004), 91500 (Wiedenbeck et al., 1995) |
| Data processing package | Iolite (Paton et al., 2010) |
| Uncertainty level and propagation | Ages are quoted at 2 sigma absolute, propagation is by quadratic addition according to Horstwood et al. (2016). Reproducibility and age uncertainty of reference material are propagated. |
| Quality control / Validation | 91500: concordia age = 1065.8±3.4 Ma (n=28; MSWD=0.94) |

References

Hiess, J.; Condon, D.J.; McLean, N.; Noble, S.R. 2012. 238U/235U systematics in terrestrial uranium-bearing minerals. Science 335: 1610-1614.

Horstwood, M.S.A.; Košler, J.; Gehrels, G.; Jackson, S.E.; McLean, N.M.; Paton, C.; Pearson, N.J.; Sircombe, K.; Sylvester, P.; Vermeesch, P.; Bowring, J.F.; Condon, D.J.; Schoene, B. 2016. Community-Derived Standards for LA-ICP-MS U-(Th-)Pb Geochronology – Uncertainty Propagation, Age Interpretation and Data Reporting. Geostandards and Geoanalytical Research 40(3): 311-332.

Jackson, S.E.; Pearson, N.J.; Griffin, W.L.; Belousova, E.A. 2004. The application of laser ablation-inductively coupled plasma-mass spectrometry to in situ U–Pb zircon geochronology. Chemical Geology 211: 47-69.

Nosenzo, F.; Manzotti, P.; Poujol, M.; Ballèvre, M.; Langlade, J. 2022. A window into an older orogenic cycle: P–T conditions and timing of the pre-Alpine history of the Dora-Maira Massif (Western Alps). Journal of Metamorphic Geology 40(4): 789-821.

Paton, C.; Woodhead, J.D.; Hellstrom, J.C.; Hergt, J.M.; Greig, A.; Maas, R. 2010. Improved laser ablation U-Pb zircon geochronology through robust downhole fractionation correction. Geochemistry, Geophysics, Geosystems 11(3): Q0AA06.

Wiedenbeck, M.; Alle, P.; Corfu, F.; Griffin, W.L.; Meier, M.; Oberli, F.; von Quadt, A.; Roddick, J.C.; Spiegel, W. 1995. Three natural zircon standards for U-Th-Pb, Lu-Hf, trace element and REE analyses. Geostandards Newsletter 19:1-23.